“Evolution as a Tool”
Bio-Hacking:
Evolution as a Tool
A short talk on a new technology
for protein engineering
and
Building a Phagestat
A longer talk on building the necessary devices
Peter Reintjes

---
Evolution as a Tool:
Use a virus to evolve a custom protein
PACE - Phage Assisted Continuous Evolution
(Esvelt et al. Nature, April 2011)
1. You need to build a PhageStat to maintain
a population of evolving virus (Husimi 1989)
Off the shelf hardware for about $30,000
- or -
DIY hardware + open source software ~$1000

---
Innatrix Motivation #1
Motivation
- Shigella kills 1,000,000 people – mostly children in the developing world – every year
- Shigella without the extracellular proteases Pic and SepA is harmless
- Hypothesis: We can evolve protease inhibitors with strong binding to Pic and SepA to diminish Shigella's virulence
- Hypothesis: These engineered proteins produced by a probiotic ( lactobacillus ) could provide inexpensive, long-term immunity

---
NCMLS Motivation
- 68% of general population believe in evolution
- 98% of scientists believe in evolution
  ...depending on your definition of 'scientist', I suppose.

---
Technology Potential
Protein-based pharmaceuticals
- Large proteins have reduced side-effect potential over small-molecule drugs which cross metabolic barriers and have more interaction potential
- Proteins can be more easily metabolized without stressing renal function
- Proteins degrade more quickly in the waste stream
- Binding affinity is the principal characteristic of metabolic processes and pharmaceuticals
- Increased binding affinity would lower dosage for any protein therapeutic

---
PACE: Phage Assisted Continuous Evolution
Evolution = Mutation + Selection
- Bacterial virus replicates inside E. coli which creates mutations by making mistakes when copying the viral DNA
- E. coli rewards improvements to evolving protein by providing a protein that increases viral infection

---
PowerPoint Presentation

---
PACE: Phage Assisted Continuous Evolution
M13 virus (phage) with the sequence to evolve replacing its fusion gene
E. coli with two extra plasmids (chromosomes)
  - Mutation: Arabinose activated mutagenesis
  (error-prone DNA polymerase)
  - Selection: M13 fusion gene activated by
  a particular protein interaction
  - E. coli host cells flow through “lagoons”
with populations of evolving viruses

---
Experiments can run from four to six days requiring reliability and automation
(1000 generations X a billion mutations)
Off-the-shelf solution requires integration as well as more hands-on oversight
(three shifts of lab personnel?)

---
Building a Phagestat
A device to maintain a stable population of bacterial phage (virus)
A Cellstat to produce (E. coli) host cells
Control of input and output flow
Inputs: Host cells and inducers
Output: Sampling and waste output

---
Building a Cellstat
To maintain a continuous supply of host cells (E. coli)
1. A Thermostat maintains cells at body temperature
2. A Turbidostat maintains cell density with nutrient dilution
3. Isolation from bacterial phage:
  Cellstat output is the Phagestat input

---
1. For PACE the flow rate must be:
1. Fast enough to force host cells through in one cell lifetime, so the only mutations are in the phage sequence
2. Slow enough to prevent washout of phage even with large variations in the viral reproduction rate

---
Design Principle #1
Replace expensive, precision equipment which is not fixable / hackable by team members with extremely simple hardware
2. Produce the required precision with computer control

---
What can we learn from Astronomy?
.
1. Kepler 22b:
2. Twice as big as Earth, in the habitable zone, could be an ocean planet, 95% confidence that it has no moons greater than ½ Earth mass
3. It is 600 light-years away !
There is a lot of information in whatever signals happen to come your way

---
Design Principle #2
We may not need to put any specialized sensors into the apparatus
2. We may be able to measure everything without any physical contact with the experimental materials

---
Raw Materials
Arduino micro-controller
Raspberry Pi / Linux
Webcam
Python programming language
OpenCV image processing software
PIR (Passive InfraRed) temp sensor
LEDs, resistors, motors, magnets
Discarded flatbed scanner
Styrofoam shipping containers
PVC plumbing hardware
3D printer

---
Turbidostat   version 1.0
Turbidostat version 1.0

---
Non-contact level-sensing
Non-contact level-sensing

---
Material Costs for a “PACE capsule” < $1000
Material Costs for a “PACE capsule”
$800
- Raspberry PI main computer ($40)
- Wide-angle USB or IP Camera ($100)
- Arduino Mega 2650 ($15)
- PIR (non contact) Temperature Sensor ($35)
- Laser, LEDs, Photo transistor ($20)
- Insulated boxes (Uline Styrofoam) ($20 X 2)
- Heating Elements ($25 X 2)
- Stirring Motors w/magnets ($5 X 5)
- Aquarium air pump ($35)
- Rotary valve motor, valves, peristaltic pumps ($150)
- Miscellaneous Hardware-PVC ($70)
- 5V and 12V Power Supplies ($20)
- Glassware ($200)
- Tubing + Nutrient (operating cost)

---
More like a 'lander' after the addition of $80 worth of PCV
More like a 'lander' after adding $80 worth of PVC

---
Hardware Projects
A multi-channel Autosampler from a flatbed scanner plus 3D printed parts
Computer controlled heater and magnetic mixer for biological samples
3D printed parts to create a 16-channel computer-controlled valve for ~$50
Low-voltage heating for liquid environments

---
Software Projects
Measuring fluid levels with OpenCV
Detecting bioluminescence with OpenCV and a USB camera
Writing an Arduino PID controller (PID = Proportional / Integral / Differential)

---
Problems:
Bioluminescence detection requires expensive high-voltage photomultiplier tubes and complex plumbing
3D printers can't make stainless steel parts
DIY electronic sensors must be sterile, but fail after being autoclaved repeatedly
The $10,000 environmental chamber failed catastrophically mid-experiment
Silicon/Tyvek tubes ruptured after several days in a constantly running peristaltic pump

---
Problem:
Bioluminescence detection requires expensive high-voltage photomultiplier tubes and complex plumbing
Solution:
In a dark room, point a heat gun at a webcam and take a picture, this is your camera's noise image
Then take 100 pictures and add the frames up using: +Green–(Red+Blue)/2 finally, subtract the noise image

---
Problem:
3D printers can't make stainless steel parts
Solution:
Perhaps you simply need parts with stainless steel edges or surfaces
Buy uncoated stainless steel welding rods $6/lb and design plastic 3D parts with openings for rods and/or plates

---
Problem:
DIY electronic sensors must be sterile, but fail after being autoclaved repeatedly
Solution:
Use non-contact sensing
PIR (passive InfraRed) Temperature sensor
Camera, laser and LED lighting, OpenCV image processing for turbidity(cloudiness), bioluminescence, pH, and level sensing

---
Problem:
The $10,000 environmental chamber failed catastrophically mid-experiment
Solution:
LM35DZ temperature sensor ($1.50)
Ardweeny ($10) running PID algorithm
1500W heat gun from Lowes ($25)
 with a solid-state relay ($10)

---
Problem:
Tyvek tubes ruptured after several days in a constantly running peristaltic pump
Solution:
Design the system for gravity/siphon flow and use peristaltic pumps for priming
 Peristaltic pumps now only run for a fraction of the total experiment duration

---
Summary of Hacks:
Styrofoam boxes for thermal isolation
Black garbage bags for optical isolation
Ardweeny controlling a heat gun
Stovetop coil + 36V printer power supply,
50 ohm 10W power resistors
 for low-voltage heating (up to 40ºC)
 Motor($1),T-nut,2 magnets = Magnetic stirrer
Camera + image processing = sensors
Pizza crisper pan, motor, 3D parts = valves
Flatbed scanner, motor, 3D parts = sampler

---
Know Wonder Moment

---
Bio-Hacking
ApE – an Open Source plasmid editor
PCR Thermocycler ($25): Ardweeny,Peltier Junction, Fan,drilled aluminum block
Gibson and Gateway assembly reagents (to assemble plasmids)
Phage Assisted Continuous Evolution
SPATEs*: The Achilles Heel of Microbial Pathogens? Tendon
*Serine Protease AutoTransporters of Enterobacteriaceae

---
Details
Minimize pressure and stress required for tube-valve

---
Bluetooth
Cellstat
System Controller
Raspberry Pi
RF Link
Phagestat
Lagoon Controller
Arduino
Phagestat
Lagoon Controller
Arduino
Phagestat
Lagoon Controller
Arduino
Phagestat
Lagoon Controller
Arduino

---
Cellstat
Nutrient
pump + valve
Bluetooth Serial
Host Cell
pump + valve
mixer
valve
Arduino
Mini Pro
Waste/Sample pump + valve
Raspberry Pi
Debian Linux
Remote Desktop
OpenCV
Python
Bluetooth
light
valve
heater
valve
camera
Communication
valve
 T sensor
magnetic mixer
Phagestat
1 of 4
aeration
lights
Bluetooth Serial
Nutrient heater Host cell heater
PIR Temp Sensor

---
Bio-Hacking:
Evolution as a Tool
A short talk on a new technology
for protein engineering
and
Building a PhageStat
A longer talk on building the necessary devices
Peter Reintjes