Evolution as a Tool
A new technology
for protein engineering
Peter Reintjes

---
Definitions
Protein
A string of amino acids that folds up into a biological machine. Life is all about proteins bumping into other proteins and causing them to do (or not do) something.
DNA
A string a nucleotides (A,T,C,G) that specifies the sequence of amino acids that make up a protein
E. coli
A group of bacteria that includes the E. coli in your stomach and other harmless types that we use in the laboratory as well as Shigella and Salmonella
Bacterial Virus = Phage
A DNA strand (rogue code) in a package that can hijack a cell and cause the cell to produce copies of the virus
-stat
A device to maintain a set of conditions

---
Evolution as a Tool:
Use a virus (phage) to evolve a custom protein
PACE - Phage Assisted Continuous Evolution
(Esvelt et al. Nature, April 2011)
You need to build a Phagestat to maintain
a population of evolving virus (Husimi 1989)
Off the shelf hardware for about $30,000
- or -
DIY hardware + open source software ~$1000

---
Motivation
Shigella kills 1,000,000 people – mostly children in the developing world – every year
Shigella without the extracellular proteases Pic and SepA is harmless
Hypothesis: We can evolve protease inhibitors with strong binding to Pic and SepA to diminish Shigella's virulence
Hypothesis: These engineered proteins produced by a probiotic ( lactobacillus ) could provide inexpensive, long-term immunity

---
Evolution = Variation + Selection
Bacterial virus replicates inside E. coli which creates variation by making mistakes when copying the viral DNA
E. coli rewards improvements to evolving protein by providing a proportional amount of a protein needed for virus to propagate

---
PACE: Phage Assisted Continuous Evolution
M13 virus (phage) with the sequence to evolve replacing its fusion gene
E. coli with two extra plasmids (chromosomes)
Variation: Arabinose activated mutagenesis
(error-prone DNA polymerase)
Selection: M13 fusion gene activated when the evolving protein binds to a target
(tighter binding causes more fusion protein to be made)
E. coli host cells flow through “lagoons”
with populations of evolving viruses

---
Experiments can run from four to six days requiring automation and reliability
(1000 generations X a billion mutations)
Create 1012 variations of the 'starting' protein
and keep the one with the tightest binding

---
Building a Phagestat
A device to maintain a stable population of bacterial phage (virus)
A Cellstat to produce (E. coli) host cells
Control of input and output flow
Inputs: Host cells and inducers
Output: Sampling and waste output

---
Building a Cellstat
To maintain a continuous supply of host cells (E. coli)
1. A Thermostat maintains cells at body temperature
2. A Turbidostat maintains cell density with nutrient dilution
3. Isolation from bacterial phage:
Cellstat output is the Phagestat input

---
For PACE the flow rate must be:
1. Fast enough to force host cells through in one cell lifetime, so the only mutations are in the phage sequence
2. Slow enough to prevent washout of phage even with large variations in the viral reproduction rate

---
Building a Phagestat
Off-the-shelf solution:
Cost ~$30,000
Requires integration of different parts
Hands-on oversight
(three shifts of lab personnel?)

---
Design Principle #1
Replace expensive, precision equipment which is not fixable / hackable by team members with extremely simple hardware
Produce the required precision with computer control

---
What can we learn from Astronomy?
.
Kepler 22b:
Twice as big as Earth, in the habitable zone, could be an ocean planet, 95% confidence that it has no moons greater than ½ Earth mass
It is 600 light-years away !
There is a lot of information in whatever signals happen to come your way

---
Design Principle #2
We may not need to put any specialized sensors into the apparatus
We may be able to measure everything without any physical contact with the experimental materials

---
Raw Materials
Arduino micro-controller
Raspberry Pi / Linux
Webcam
Python programming language
OpenCV image processing software
PIR (Passive InfraRed) temp sensor
LEDs, resistors, motors, magnets
Discarded flatbed scanner
Styrofoam shipping containers
PVC plumbing hardware
3D printer

---
Bluetooth
To get rid of all the wiring

---
Turbidostat version 1.0
$25
vs.
$10,000

---
Non-contact level-sensing

---
Material Costs for a “PACE capsule”
Raspberry PI main computer ($40)
Wide-angle USB or IP Camera ($100)
Arduino Mega 2650 ($15)
PIR (non contact) Temperature Sensor ($35)
Laser, LEDs, Photo transistor ($20)
Insulated boxes (Uline Styrofoam) ($20 X 2)
Heating Elements ($25 X 2)
Stirring Motors w/magnets ($5 X 5)
Aquarium air pump ($35)
Rotary valve motor, valves, peristaltic pumps ($150)
Miscellaneous Hardware-PVC ($70)
5V and 12V Power Supplies ($20)
Glassware ($200)
Tubing + Nutrient (operating cost)
$800

---
More like a 'lander' after adding $80 worth of PVC

---
Problems:
The $10,000 environmental chamber failed catastrophically mid-experiment
Bioluminescence detection requires expensive high-voltage photomultiplier tubes and complex plumbing
3D printers can't make stainless steel parts
DIY electronic sensors must be sterile, but fail after being autoclaved repeatedly
Silicon/Tyvek tubes ruptured after several days in a constantly running peristaltic pump

---
Problem:
The $10,000 environmental chamber failed catastrophically mid-experiment
Solution:
LM35DZ temperature sensor ($1.50)
Ardweeny ($10) running PID algorithm
1500W heat gun from Lowes ($25)
with a solid-state relay ($10)

---
Problem:
Bioluminescence detection requires expensive high-voltage photomultiplier tubes and complex plumbing
Solution:
In a dark room, point a heat gun at a webcam and take a picture, this is your camera's noise image
Then take 100 pictures and add the frames up using: +Green–(Red+Blue)/2 finally, subtract the noise image

---
Problem:
3D printers can't make stainless steel parts
Solution:
Perhaps you simply need parts with stainless steel edges or surfaces
Buy uncoated stainless steel welding rods $6/lb and design plastic 3D parts with openings for rods and/or plates

---
Problem:
DIY electronic sensors must be sterile, but fail after being autoclaved repeatedly
Solution:
Use non-contact sensing
PIR (passive InfraRed) Temperature sensor
Camera, laser and LED lighting, OpenCV image processing for turbidity(cloudiness), bioluminescence, pH, and level sensing

---
Problem:
Tyvek tubes ruptured after several days in a constantly running peristaltic pump
Solution:
Design the system for gravity/siphon flow and use peristaltic pumps for priming
Peristaltic pumps now only run for a fraction of the total experiment duration

---
Summary of Hacks:
Styrofoam boxes for thermal isolation
Black garbage bags for optical isolation
Ardweeny controlling a heat gun
Stovetop coil + 36V printer power supply,
50 ohm 10W power resistors
for low-voltage heating (up to 40ºC)
Motor($1),T-nut,2 magnets = Magnetic stirrer
Camera + image processing = sensors
Pizza crisper pan, motor, 3D parts = valves
Flatbed scanner, motor, 3D parts = sampler

---
Know Wonder Moment

---
Protein-based pharmaceuticals
Reduced side-effect potential
Proteins are easily metabolized
Proteins degrade quickly in the waste stream
Binding affinity is the principal characteristic of metabolic processes and pharmaceuticals
Increased binding affinity lowers dosage

---
Evolution as a Tool
A new technology
for protein engineering
Peter Reintjes

---
Bio-Hacking
ApE – an Open Source plasmid editor
PCR Thermocycler ($25): Ardweeny,Peltier Junction, Fan,drilled aluminum block
Gibson and Gateway assembly reagents (to assemble plasmids)
Phage Assisted Continuous Evolution
SPATEs*: The Achilles Heel of Microbial Pathogens? Tendon
*Serine Protease AutoTransporters of Enterobacteriaceae

---
Details
Minimize pressure and stress required for tube-valve

---
Bio-Hacking:
Evolution as a Tool
A short talk on a new technology
for protein engineering
and
Building a PhageStat
A longer talk on building the necessary devices
Peter Reintjes