An Inexpensive Phagestat for Continuous Evolution Research

Table of Contents

# Abstract
Continuous evolution systems (PACE, PATHE) provide a powerful approach to protein engineering to increase Protein-DNA and Protein-Protein binding affinities and selectivity through directed evolution. Ad-hoc systems built from commercially available equipment have material costs on the order of $30,000 as well as high labor costs because of the requirement of continuous monitoring. Alternatively, by using inexpensive computational systems, commodity electronic and plumbing, and a few custom 3D-printed components, a bio-reactor can be built for approximately $1000 which reduces oversight requirements through non-contact sensing, image processing, and computer-controlled feedback. This inexpensive and easy-to-use apparatus will be a scalable platform for using directed evolution for protein engineering.
# Problem
Protein Engineering consists principally of finding DNA sequence modifications that result in a desired protein structure. The difficulty of predicting precise three-dimensional shape and charge distribution from a sequence change means that traditional engineering approaches are often impractical.
And even if a precise correspondence between sequence changes and conformational changes were known, it would still require precise knowledge of the target's shape and charge distribution in vivo, which might be significantly different from the structural information available through crystallography or NMR. However, when the protein engineering task is to create a protein with stronger and more selective binding to a particular target, we can use evolution to perform a hill-climbing search among randomly generated variations of a protein.

Directed evolution is a method whereby in vivo interaction between a target protein and mutable protein can be used as a selection mechanism to discover novel sequences with desired properties. Given a target and a provisional protein which have some binding affinity, directed evolution is a method of finding sequence changes which increase binding and selectivity [PACE].
We give the name “phagestat” to an apparatus which maintains a population of bacterial virus undergoing directed evolution through interaction with a bacterial host that has been programmed to mutate, select, and propagate that virus [PHAGESTAT].
An inexpensive, reliable, and highly automated phagestat will allow many more directed evolution experiments to be performed. Although the task of designing plasmids for the host cell often takes longer than running the evolution process, the inexpensive phagestat opens the possibility of performing many parallel experiments with different starting sequences and sequentially evolving sequences for increased binding, selectivity, or negative selection pressure for unwanted properties (Dickson).
# Current Status of Research
Prototypes of all individual components, including the turbidostat, host cell incubator (cellstat), computer vision level monitoring, leakage alarm, and automatic sampling system have been built and tested. Reliability testing over several weeks has been performed on various subsystems including the heating/mixing/aeration cell culture system, phage lagoon mixing, heating, and lighting systems as well as communications between these subsystems (cellstat and lagoons) and the main computer. Platform independent control software, written in Python and Prolog, and using OpenCV for image processing has been tested on Microsoft Windows (XP, 7) and Linux (Debian8, Ubuntu, and Raspian).
We know of no similar work to reduce the cost of constructing and operating phage-based continuous evolution environments. We acknowledge that the capital and labor costs of running such experiments may be of minor importance to practitioners of academic biological research. Researchers currently involved in active research in this area probably do not regard a $30,000 capital cost and oversight requiring several lab technicians as a primary concern. However, the effort that goes into setting up and debugging these systems has not being fully quantified, and we believe that streamlining the mechanics of maintaining an evolving phage population will free these practitioners to concentrate on plasmid design and conduct many more experiments within their time and budget constraints.
This work requires expertise in areas outside of molecular biology, and its goal is to lower the cost of entry into this area of research rather than to achieve any particular protein engineering task. That said, the current design is driven almost entirely by the requirements of ongoing work by Drs. Edgell, Charles, and Collier to enhance binding for protease inhibitors of bacterial extra-cellular virulence factors.

## Technical Goals

Develop software to assist calibration of the EvoStat using external references for luminosity and cell density, thus exploiting shared equipment while keeping the cost of each phagestat low.

Evaluate the use of several commodity USB and WiFi camera systems to accommodate hardware cost or performance improvements and availability.

Test operation over many hours of unattended operation. Goals include reliable operation for at least 100 hours (four days) and rapid detection and communication of problems (e.g., controlled suspension of operation and sending of text messages after leak detection).
Web-based interface for viewing data logs, changing parameters such as selection pressure modulation (inducer schedule), and live images from the main camera and microscopic/bio-luminescence camera.
Develop procedures for replacing any subsystem without interrupting the long-term experiment.

For example, the most difficult of these operations would involve cellstat replacement. First, a second cellstat would be filled with nutrient, brought to temperature, and inoculated from the current cellstat. Calibration data are supplied as the culture is brought to early log phase growth rate. Once achieved, the liquid flow through the system will be paused and the new cellstat swapped into the system. Restarting the system with the new cellstat name in the configuration file will complete the cellstat change. Each instance of cellstat, lagoon, or auto-sampler is given a unique name which will appear during Bluetooth discovery. The device names which appear in a configuration file define the components in that system. Auto-sampler replacement merely requires the controlled interruption and resumption of normal flow operations, while lagoon replacement will require a greatly reduced flow rate and possibly extra inducers to re-establish the viral population.
# Project Plan

Three complete EvoStat systems will be built: One, nearing completion as of this writing, to support the work of Edgell et al. at the Genome Sciences building on the campus of UNC-Chapel Hill. A second will be constructed at the local maker space in Durham, NC (Splatspace) where we have access to 3D printers, an electronic workbench, woodworking, and machine shop, and a third system at the Museum of Life and Science, where principal construction of components for all three systems will occur.
Each EvoStat will consist of a tower frame, main computer, one cellstat, four lagoons, and one auto-sampler. An additional cellstat and an auto-sampler along with a few additional lagoons may be constructed as backup units. We have no plans to run active cultures at the Museum of Life and Science or the Makerspace. When the design stabilizes and the second and third EvoStat have demonstrated reliability, they will be moved to UNC-CH or another laboratory to be used for evolution experiments.

Patent application(s) will be completed covering various aspects of the project including the design of three custom components to perform turbidometer and auto-sampling, and software components for process control, communication, and image processing.
Improve the design to exploit low-cost alternatives to specialized equipment, and allow for the construction and substantial modification of the device to be performed by the laboratory personnel who will be using it.
Describe the studies that will be performed to meet the stated technological goals, including key technical milestones and a timeline narrative.

Aims		Months 1-3	Months 4-6	Months 7-12
Level detection
Cell density
Luminosity
Auto-calibration

Table 1 Project Timeline

Support will provide funding for two days per week PI for six months: one day per week at the Museum of Life and Science for primary construction and software development of three complete EvoStats and one day per week on-site at the Genome Sciences building (biology department) at UNC-Chapel Hill integrating and maintaining the EvoStat for use by researchers.

Completed design and construction of major sub-systems:
1. EvoStat Frame
  1. Three-level PVC support platform
  2. Insulated dark environmental chamber
  3. Camera
  4. Power supplies: 5-, 12-, and 36-volt
Cellstat
  1. Turbidostat
  2. Digital thermometer
  3. Heating element
  4. Meniscus light
  5. Magnetic mixer/Aeration
  6. Micro-controller + software
  7. Bluetooth communication module
2. Lagoon
  1. Thermometer
  2. Heating elements
  3. Meniscus light
  4. Magnetic mixer
  5. Pinch valve
  6. Microcontroller + software
  7. Bluetooth communication module
3. Autosampler
  1. Rolling multi-valve
  2. Sample plate positioning platform
  3. Micro-controller, stepper, motor and sample-timing software
  4. Bluetooth Communication module
4. Luminometer
  1. USB microscopic camera
  2. Dark enclosure for camera
  3. Cooling unit for camera

Deliverables:

1. Manifest of parts costing less than $1000
2. Instructions for construction, maintenance, and modification of the device

   We intend to use the devices to recapitulate the experiments done by Esvelt, Carlson, Lui, Dickson, and others at Harvard and to support the ongoing work being done by Dr. Edgell et al.
# Evaluation

### Criteria for a Go/No-Go Decision, Including Metrics

- Completed construction of three fully functional EvoStats with a parts cost of less than $2000.
- All instances capable of maintaining environmental conditions continuously for four days.
- Demonstrated ability to replace each of the major subsystems: cellstat, lagoon, and autosampler while maintaining a live culture.
- Perform several complete evolution experiments on evolving phage.

## Discuss challenges
Image processing must be adapted to constantly changing hardware and laboratory conditions.

Active cooling must be added to low-cost USB cameras to avoid the expense of a cryogenic camera (e.g. $25,000 from Hammumatsu).

Equipment must be modular and easy to move due to space restrictions.

# Commercial Proposal Narrative
  (Commercialization Advisor should participate in the preparation of this section of the proposal)
  Contact query sent to Mike Carnes, SBTDC
## Problem
Phage-based continuous evolution is a new technology with few practitioners.

## Commercial Goals

We want to encourage researchers with limited budgets to use these design ideas to perform research, possibly by requiring no licensing fee. This approach could result in the sharing of improvements to the design from an open community. We would then require licensing only for commercial exploitation of sequences resulting from the use of this device.
As evolution-based protein modification matures into an industrial process, we expect to license this design to one or more manufacturers of laboratory equipment. The device itself is a tool for sequence discovery and does not produce therapeutic proteins or chemicals so it should not be subject to any medical equipment or FDA requirements.
One company (Innatrix) is already considering a business model to employ a number of these machines and a staff with expertise in selection-plasmid design to provide contract protein optimization. Such a service would allow a drug development company with a protein therapeutic to contract out the task of affinity or selectivity optimization without the need to build and maintain phage-stats or develop in-house expertise in directed evolution. This business model is of particular interest as it benefits most directly from lower construction and operating costs of the phagestat.

There is a consultation business opportunity for introducing this technology to companies in the pharmaceutical sector.

## Market

Phage-based continuous evolution is a fairly recent development, but given the advantages of protein-based therapeutics, it seems likely that this method of protein modification will come into widespread use in pharmaceutical development.
Manufacturers of liquid handling laboratory equipment such as Tecan, Aventics, Humphrey, and Nordson EFD are potential manufacturers of the EvoStat. Others include Becton Dickinson, Topac, Intellitech, and Pope Scientific.
Potential customers include:
- Pharmaceutical research organizations engaged in protein engineering.
- Non-pharmaceutical protein engineering such as materials science.
- Protein optimization services for organizations that do not want to develop in-house phage/evolution capabilities. A protein optimization service would accept sequences for binding target and initial candidate proteins, design selection-plasmids and hybrid proteins, and perform the evolution process.

## Intellectual Property

The PACE (Phage Assisted Continuous Evolution) technique is the subject of patents filed by Harvard University (ca. 2011). Modifications to this process, such as PATHE (Edgell et al. unpublished correspondence) represent improvements to this technique which themselves may or may not be constrained by the PACE patents. These patents relate to the molecular biology, phage selection, the use of evolving protein as a proxy for phage fusion protein, and other matters relating to the molecular biology of mutation and selection. This IP does not appear to impact phagestat design and our improvements are based on well-known and largely open technologies. A description of the prior art can be found in Husimi (1989).
As far as we know, none of the details of operation of this design are subject to any IP constraints.
Hardware designs and image processing software for flow-control, turbidity sensing, and detection of bioluminescence. A novel low-cost non-contact turbidometer design. Integrated communications architecture to reduce wiring and simplify set-up of equipment. Techniques for repair and replacement of subsystems during operation. Self-calibrating and continuous monitoring of all system parameters.
Data logging and alerts utilizing Internet capabilities.
( overview of the project, description of technology, Commercialization Advisor, summary of the technical and commercialization goals)
# Technical Proposal Narrative

Describe the work that has been done to date, both in the lab of the scientist and in the field.
The following is a temporary table listing different apects and subsystems of the current design
Devices	Measurements	Software
Heaters	Temperature	User Interface
Photo-transistor	Turbidity	Micro-controllers
Camera 1	Level detection	Image Processing
Camera 2	Luminosity	Image Processing
Valves	Leak Detection	Notification (email/text)
	Flow control	Auto-calibration

Mixing
Aeration
Commuications	Data Logging	Bluetooth, Web server, database

## Evaluation
        1. The completed design should have a parts cost on the order of $1000 and can be constructed in a few weeks by an electronic technician.
        2. Construction should require a minimal tool set, perhaps consisting of a drill press, hand drill, table saw, chop/miter saw, deburring tool, and thread tap.
        3. Demonstration that the device can perform long term evolution experiments with the claimed reduction in oversight.
  2. Achieving sufficient image sensitivity for bio-luminescence through the active cooling of commodity USB cameras (Logi-Tech, Dino-Lite USB microscope).
  3. Evalutate potential markets for a Go/No-go decision about patent filing.
## Patent protection
Ultimately, many practitioners may want to perform the continuous evolution process with requirements different from academic research. We see an opportunity to create a commercial version of the EvoStat with an emphasis on using directed evolution as an industrial process.
## Licensing

## References

[PhageStat] Selection and Evolution of Bacteriophages in Cellstat, Yuruzu Husimi, Advances in Biophysics, Vol. 25, pp.1-43 (1989).
# [PACE] A system for the continuous directed evolution of biomolecules,Esvelt KM, Carlson JC, Liu DR, Nature. 2011 Apr 28;472(7344):499-503. doi: 10.1038/nature09929. Epub 2011 Apr 10.

[PATHE] Dr. Marshall Hall Edgell, unpublished private communication.